Aqua regia acid digestion of soil for major and trace elements

WSW/003/27.6.2016

2.2

Rothamsted Research Analytical Unit

Protocol and procedure manual

Principle

Digestion of soils with aqua regia acid is used for the determination of major and trace elements. The samples are digested in tubes, which are heated in time- and temperature-controlled heating blocks, filtered and analysed by ICP.

This method agrees well with the reflux method but is easier to operate and requires less attention. Although this method does not give absolute total contents, it does give results sufficiently close to accepted values for different types of soils and sludges for it to be of value in the routine monitoring of metal contents.

The method can be modified to suit different types of soil, eg to overcome excessive frothing when digesting highly calcareous soils, extra hydrochloric (HCl) acid is added before the aqua regia and heating stages. Peaty soils should be ashed at 450°C before digestion for three hours.

Soils for metal analysis MUST ALWAYS BE GROUND USING AN AGATE MILL, not a metal one.

Apparatus

1	Carbolite heating block with time and temperature controller.

2	Digestion tubes, 150 x 22 mm, graduated to 25 ml, with stopper.

- Optifix variable volume 0–10 ml dispenser (one for water and one for HCl)

- Gilson 5 ml automatic pipette (set to dispense 1 ml nitric acid)

5	Vortex tube mixer (whirlimixer)

6	Polythene vials – 25 ml (Sterilin)

7	Scrubbed fume cupboard

Reagents

1	HCl AR (SG 1.18 g/ml, 37 per cent)

2	Nitric acid (Primar or Aristar) (SG 1.42 g/ml, 70 per cent)

3	25 per cent HCl: add 250 ml HCl AR to water, mix and make up to 1,000 ml with 18 MΩ deionised water

Procedure

1	Weigh out 0.250 g of finely milled, air-dried soil into a weighing boat and transfer to a 25 ml graduated digestion tube (use aqua regia digestion tubes only).

- To each tube, add 4 ml HCl (dispenser) and 1 ml nitric acid (Gilson pipette) (ie 5 ml aqua regia), whirlimix and leave to stand at room temperature for at least two hours. Add around the side of the tube ensuring all soil is washed down into the tube (do not premix the aqua regia).

See notes below for method for highly calcareous soils.

- Whirlimix again, place in the Carbolite heating block (54 hole) and heat using the following programme:

| Ramp no | PR temp. rise °C/min | PL dwell temp. °C | PD dwell time mins |
| --- | --- | --- | --- |
| 1 | 2 | 25 | 120 |
| 2 | 2 | 60 | 180 |
| 3 | 2 | 105 | 60 |
| 4 | 2 | 125 | 120 |
| 5 | Step | 50 | End |

It may be necessary to whirlimx occasionally at or below 60°C to prevent a cap of partly digested material rising up the tube.

The above heating regime is carried out overnight. The tubes should be cool by morning.

4	Using a dispenser, add 5 ml 25 per cent HCl to each tube, whirlimix and reheat at 80°C for one hour.

5	Whirlimix, add approximately 18 ml of deionised water (dispense two 9 ml aliquots) and reheat at 80°C for 30 mins. Remove from the block and allow to cool before making up to volume. Stopper and mix well by shaking (invert fully).

6	Immediately filter the contents of the tube through a Whatman no 40 filter paper into a Sterilin vial. Discard the first 5 ml of filtrate.

7	Blanks and appropriate in-house standard materials should be included in each block digested.

8	Duplicate sample digestions of at least one in every ten samples should be included in each batch.

9	Samples should be presented for analysis by ICP.

Washing Procedures

- Digestion tubes should be rinsed in deionised water, cleaned using a plastic-coated bottle brush, rinsed again and then placed in a 5 per cent HCl acid bath overnight.

- Tubes are then placed in a deionised water bath.

- Tubes are then rinsed six times in deionised water and dried in a drying cabinet.

- Stoppers should be rinsed twice with deionised water then acid washed overnight in 5 per cent HCl.

- Stoppers are then rinsed six times in deionised water and dried in a drying cabinet.

- Polythene funnels should be rinsed in deionised water then placed in a 5 per cent HCl bath for approximately one hour.

- Funnels are then placed in a deionised water bath.

- Funnels are then rinsed six times in deionised water and dried in a drying cabinet.

Notes

1	Read the Control of Substances Hazardous to Health Regulations (2002) forms relating to the preparation of samples by aqua regia digestion before using this method. Use suitable protective clothing (spectacles and gloves) (see Docport: RR_QA_Doc/1021/001).

2	When working with calcareous soils, add 2 ml HCl and whirlimix to reduce frothing, add a further 2 ml of HCl and whirlimix. When the frothing has ceased, add 1 ml of HCl and 1 ml nitric acid. The additional 1 ml of HCl makes up for the acid used in reducing the calcium carbonate.

3	Peaty soils with high organic matter should be weighed into a silica crucible, placed in a cold muffle furnace and ashed at 450°C for three hours. When cool, the ash should be washed into a graduated 25 ml tube with 8 ml of 50 per cent HCl followed by 1 ml nitric acid. Then proceed from step 3. Under these conditions, copper may be lost by volatilisation during ashing.

4	Refer to the quality assurance manual in room 105 for quality control details relating to sample preparation, internal and external standard preparation and selection and procedure protocols.

References

1	S P MCGRATH & C H CUNLIFFE, A Simplified Method for the Extraction of the metals Fe, Zn, Cu, Ni, Cd, Pb, Cr, Co and Mn from Soils and Sewage Sludges. J. Sci. Food Agric. 1985.36 794–798.

2	T T GORSUCH, The Destruction of Organic Matter. Pergamon Press, Oxford 1970