Sodium bicarbonate extraction of soil (Olsen) for available P

determination by Skalar continuous flow analysis

WSW/003/27.6.2016

2.5

Rothamsted Research analytical unit

protocol and procedure manual

Principle

Phosphate ions (HPO4 and H2PO4) are present in soils in association with calcium, aluminium and ferric ions. If the concentration of these metallic ions is reduced, the concentration of phosphate ions increases to maintain the various solubility products at their constant values.

An alkaline (pH 8.5) bicarbonate solution can repress the concentration of calcium ions by precipitation as calcium carbonate and of aluminium and ferric ions by precipitation as hydroxides. Thus, phosphate ion concentrations are increased and available phosphate can be extracted from soil by shaking with alkaline sodium bicarbonate. The extraction is dependent on time of contact and temperature.

A disadvantage of this extraction procedure in dissolving organic matter from humus-rich soils can be overcome by treating the solution with active charcoal during the shaking or after filtration.

Phosphorus in the bicarbonate solution is determined by the phospho-molybdenum blue method on the Skalar SANPLUS System (continuous colourimetric flow analysis).

Apparatus

1	Temperature-controlled orbital shaker

- Polythene bottles 100 ml

- Polythene filter funnels

- 250 ml glass conical flasks

- Fluted Whatman no 42 filter papers (150 mm)

Analytical instrumentation

Skalar SANPLUS System (continuous colourimetric flow analysis)

- Skalar Analytical BV, Breda, Netherlands (manufacturer)

- Skalar (UK) Ltd, 8 Warren Yard, Warren Park, Milton Keynes, Buckinghamshire (British supplier)

Reagents

1	0.5 M sodium bicarbonate solution: dissolve 42 g sodium bicarbonate (NaHCO3) and 0.72 g sodium hydroxide pellets in 900 ml 18 MΩ deionised water. Adjust the pH to 8.5 using a saturated solution of sodium hydroxide or concentrated sulphuric acid and make up to 1 litre with water. Mix thoroughly (210 g NaHCO3 / 3.6 g NaOH in 5 litres) – store in a room at 20°C

2	Activated charcoal

Procedure (to be carried out in a 20ºC laboratory)

1	Weigh 5 g of air-dried soil, ground to pass a 2 mm sieve, in a weighing boat and transfer to a 250 ml conical flask.

2	Add approximately 0.1 g activated charcoal (tip of a spatula). Cover the flasks.

3	Add 100 ml 0.5 M sodium bicarbonate, reseal, swirl the flask and immediately place on the shaker. THESE STEPS MUST BE CARRIED OUT SPEEDILY TO AVOID DELAYS WHICH MAY VARY THE CONTACT TIME.

4	Shake the flask in a temperature-controlled cabinet orbital shaker at 20°C for exactly 30 mins at exactly 120 RPM per minute (40 flasks).

5	Swirl the flask (to mix the extract and soil) and filter immediately through a Whatman no 42 filter paper, discarding the first 10 ml of filtrate. Remove the filter paper and whatever solution remains in the filter paper, exactly 45 minutes after discarding the first 10 ml. BETWEEN RUNS, IT IS ESSENTIAL TO KEEP THE CONTACT TIME CONSTANT.

6	After filtration, the sample should be analysed immediately. DO NOT STORE THE FILTRATE.

7	If the extracts are highly coloured, shake the filtered extract with 1 g of activated charcoal and filter through a Whatman no 42 filter paper.

8	Bottles should be capped, placed in a tray and taken to the anion laboratory with an analysis information card ready for immediate analysis by Skalar continuous flow.

9	Blanks and in-house standard materials should be included in each batch shaken.

10	Duplicate extractions of at least one in every 10 samples should also be included.

11	Sample extracts should be analysed within 2 hours of extraction by Skalar continuous flow analyser.

Procedure modifications when extracting ISE samples

- Run a set of soil samples in duplicate using the volume method:
Make sure the soil is well mixed. Take a 2.5 ml scoop of soil, tap gently once, smooth off the excess soil using a flat surface (the side of a spatula is fine), tip into the flask and repeat to make 5 ml of soil. Continue as for the standard method but filter through no 2 Whatman filter paper.

- Run a second set of single soil samples using the weighing method (5 g) and following the standard test procedure using no 42 Whatman filter paper.

Procedure modification when extracting REF LAB/ISDA samples

- For every Olsen P run on a REF LAB/ISDA sample, include the PBI soil standard.

Washing procedures

Conical flasks, funnels and plastic bottles:

- Should be washed twice in tap water prior to the acid bath and four times in deionised water.

- Should be acid washed fortnightly in 5% HCl overnight and rinsed as above.

OR

- Wash in the Miele lab washer (room 324) on programme 2.

Notes

1	Read the Control of Substances Hazardous to Health Regulations 2002 (COSHH) forms relating to the preparation of samples by ammonium acetate extraction before using this method. Use suitable protective clothing, ie spectacles and gloves (see Doc Port: RR_QA_Doc/1005/001).

2	Refer to the analytical section quality assurance manual in Bawden room 335 for quality control details relating to sample preparation, internal and external standard preparation, and selection and procedure protocols.

3	Activated charcoal currently in use: General purpose acid washed charcoal, Activated, Fisher Scientific C/4080/50.

References

1	S R OLSEN et al., USDA Circ

ular 939 1954