pH measurement of soil in water

WSW/001/19.6.2007

2.6

Rothamsted Research analytical unit

protocol and procedure manual

Principle

On drying, the pH value of a soil is probably modified. To estimate the true pH of the soil in equilibrium with soil water, the dried and milled soil is wetted with water.

The soil suspension used has a soil to water ratio of 1:2.5.

Apparatus

1	pH meter, temperature compensated (Jenway)

2	Combination pH electrode

3	Glass bottles with screw-top lids (60 ml)

4	Variable volume 0–30 ml dispenser

Reagents

1	Buffer solutions, pH 4.0 and pH 7.0 at 20°C (Whatman high resolution pH buffer solution) (two bottles, one for preconditioning and one for calibration)

2	Check buffer solution of pH 6

3	18 MΩ Deionised water, freshly boiled and cooled

Procedure

1	Weigh 10 g of air-dried soil (ground to pass 2 mm sieve) and transfer to a 60 ml glass bottle.

2	Using a dispenser, add 25 ml of boiled, cooled deionised water, replace lid and shake, leave to stand for 30 minutes, shake once again and leave for a further 30 minutes (total 1 hour).

For soils from the classical experiments (AMD) please see below.

3	Calibrate the pH meter with the pH 7.0 and 4.0 buffers.

- Switch on the pH meter and leave for 5 minutes for the temperature to stabilise.

- Rinse the electrode with deionised water, blot off excess water and place electrode and temperature probe into the preconditioning pH 7.0 buffer, remove and place directly into the calibration pH 7.0.

- Allow the meter reading to stabilise and press “Calibrate”. Rinse the electrode and repeat previous two steps with pH 4.0 buffer (preconditioning and calibration solutions).

- Check buffer solution 6.0. The reading should be within the expected range.

4	Shake the sample solutions again and measure immediately. Lower the electrode and temperature probe into the sample supernatant, wait 30 seconds and note the pH reading when the meter has stabilised. Remove the electrode from the sample and rinse with deionised water, blotting the excess water.

- Check the calibration with pH 7.0 after every 10 samples. If pH 7.0 varies by more than ± 0.1–0.15, recalibrate the meter.

- Reread every 10th sample at the end of the batch to check the pH values.

Soils from the classical experiments (AMD)

Any samples received from the classical experiments should be processed differently to above. Take two pH readings, one after 30 minutes and a second 90 minutes later (total 2 hours). A mean value should be calculated, but both sets of data should be returned to the sponsor.

Notes

1	Duplicate samples of one in 10 should be included in every batch (eg 2R, 12R).

2	Two in-house standard soil samples should be included in every batch, one at the beginning and one at the end.

3	Care of the rapid renew electrode. The electrode should be stored in an electrode storage solution (if unavailable, use 200 ml pH 4.0 + 1 g potassium chloride). During a single day, the electrode may be left in deionised water between batches. The electrode should be filled with 3.5 M potassium chloride solution.